Gastec Tube Datasheet

| Acetaldehyde | CH3CHO |  | NO.GAS92M |
| --- | --- | --- | --- |

Performance

| Measuring Range | 2.5 to 5 ppm | 5 to 100 ppm |
| --- | --- | --- |
| Number of Pump Strokes | 2 | 1 |
| Correction Factor | 1/2 | 1 |
| Sampling Time | 1.5 minutes per pump stroke |  |
| Detecting Limit | 0.25 ppm (n=2) |  |
| Colour Change | Yellow Red |  |
| Reaction Principle | 3CH3CHO + (NH2OH)3H3PO4 H3PO4 + Base Phosphate |  |
| Coefficient of Variation | 10% (for 5 to 20 ppm), 5% (for 20 to 100 ppm) |  |
| Shelf Life | Up to 2 Years |  |
| Corrections for temperature & humidity | Unnecessary |  |
| Store the tubes in the refrigerator to keep at 10oC (50oF) or below. |  |  |

| Possible coexisting substances and their interferences |  |
| --- | --- |

| Substance | Concentration | Interference | Change colour by itself |
| --- | --- | --- | --- |
| Ammonia | - | Minus error | No discoloration |
| Acrolein | ≥3/5 time | Plus error | Produces red stain |
| Acetone | ≥2 times | Plus error | Produces red stain |
| Methyl ethyl ketone | ≥6 times | Plus error | Produces red stain |
| Methyl isobutyl ketone | ≥10 times | Plus error | Produces red stain |

| Calibration gas generation | Diffusion tube method |
| --- | --- |